Human Shoulder Finite Element Model ReadMe

Created by Sara Sadeqi

Last modified March 22, 2022

**Background**

A finite element model of the male human shoulder was created. The model is a first step in developing a validated computational framework for researching the biomechanics of the human shoulder. The model is provided, here, in its entirety. The included inp and cae files (Male-Shoulder.inp, and OpenSourseShoulderFE.cae) contain the finite element model, as constructed in Abaqus/Explicit 6.14-5. The model simulates abduction of the shoulder joint in the scapular plane. Simulation outputs include contact forces and stresses.

**Licensing**

CC0 1.0 Universal (CC0 1.0)

Public Domain Dedication

https://creativecommons.org/publicdomain/zero/1.0/

This is a human-readable summary of the Legal Code (read the full legal code here):

https://creativecommons.org/publicdomain/zero/1.0/legalcode.

No Copyright

This license is acceptable for Free Cultural Works.

Output data are provided for comparing to outputs generated by simulations on other hardware/software systems to ensure functionality of the model.

**Model Components**

The model (Figure 1) includes 3D geometries meshed with C3D10 elements for the following tissues: Humerus, Scapula, Clavicle, Humeral cartilage, Glenoid cartilage, Labrum, and Acromioclavicular ligament.

2D connector elements with tension loads are used for muscle-tendon units of the following muscles: rotator cuff (Supraspinatus, Infraspinatus, Subscapularis, and Teres minor), Mid-Deltoid, and clavicular head of the Pectoralis Major.




Figure 1) (a) Model of the intact shoulder in Abaqus displaying the different bones (humerus, clavicle, scapula, and glenoid) and soft tissue (humeral and glenoid cartilages, AC ligament, and labrum) components. (b) Meshed view of the model depicting interactions and muscle 2D connectors.

**Component Geometry**

Humerus, Scapula, Clavicle, labrum, and Acromioclavicular ligament were manually segmented from CT scans of the visible human project male subject in Mimics 19. The glenoid cartilage and humeral cartilage were made by offsetting the mesh layers in Abaqus to reach the thickness of these cartilages as mentioned in the literature [1, 2].

**Material Properties**

Material properties of the model (Table 1) include linear elastic for the bones (humerus, scapula, and clavicle) and Neo-Hookean hyper-elastic for the soft tissue (Glenoid cartilage, humeral cartilage, labrum, and Acromioclavicular ligament). Portions of the bones away from the areas of interest (humeral shaft, scapula body, and medial clavicle) were modeled as rigid bodies for computational efficiency. Values from the literature were applied to the all the model components.

Table 1) Material properties of different components of the model.

|  |  |  |  |
| --- | --- | --- | --- |
| Anatomy | Material Type | Parameters | Reference |
| Scapula | Linear Elastic | E=16 GPa, ν=0.3 | [3] |
| Humerus (Cortical Bone) | Linear Elastic | E=12 GPa, ν=0.3 | [4] |
| Humerus (Cancellous Bone) | Linear Elastic | E=250 MPa, ν=0.3 | [4] |
| Clavicle | Linear Elastic | E=17 GPa, ν=0.3 | [5] |
| Glenoid | Linear Elastic | E=1.4 GPa, ν=0.3 | [6] |
| Acromioclavicular (AC) Ligament | Hyper-elastic | C10=1.125 MPa, D1=0.19 | [7] |
| Glenoid Cartilage | Hyper-elastic | C10=1.79 MPa, D1=0.12 | [8, 9] |
| Humeral Cartilage | Hyper-elastic | C10=1.79 MPa, D1=0.12 | [8, 9] |
| Labrum | Hyper-elastic | C10=12.5 MPa, D1=0.017 | [10] |

**Interactions**

Surface-to-surface interactions with finite sliding (tangential behavior: penalty, coefficient of friction=0.01, normal behavior: “Hard” contact pressure-overclosure) were applied to the contacts between humeral and glenoid cartilages and between humeral cartilage and labrum. Tie constraints were used for connecting humeral cartilage to the humeral head bone, glenoid cartilage to the glenoid bone, and the ends of the acromioclavicular ligament to the clavicle and acromion.

**Steps, Loads, and Boundary Conditions**

The model consists of two Dynamic, Explicit steps. In the first step, a compressive displacement boundary condition (BC) of 0.8 mm was applied in the z direction of the glenohumeral joint axis (LOADAXIS) to bring the humeral and glenoid cartilages in contact. In the second step, 100о abduction rotational BC was applied over the y axis of the LOADAXIS coordinate system. Other degrees of freedom in the model were fixed. Muscle forces were also applied in the second step via connector force with amplitudes from the literature. The abduction was used to validate the model outputs against in vivo experiments [11, 12] and compare with other simulations [13-17] which modeled abduction.

**Mass Scaling**

Mass scaling may be used in Abaqus/Explicit analyses to increase computational efficiency with the potential consequence of decreased accuracy. To choose a proper mass scaling for this model, the energy outputs were monitored to make sure they satisfy the energy balance criteria mentioned in the Abaqus documentation (Abaqus 6.14 user manual, sec. 6.3.3).

**Outputs**

The main output of the model is glenohumeral contact force. However, other outputs such as stress, strain, contact area, and contact pressure can also be extracted from the model.

**Verification**

Simulation outputs are provided in the added spreadsheet. Users are encouraged to compare these outputs to those generated on their software/hardware.

**Runtime**

The time for the simulation to complete was approximately 15, 17.5, 22, and 23 hours for 84cpus-single precision, 8cpus-single precision, 84cpus-double precision analysis, and 84cpus-double precision analysis and packager, respectively.

**Additional Information**

Full details on the development of this model are published in Annals of Biomedical Engineering: <https://link.springer.com/article/10.1007/s10439-022-03018-8>.

**References**

[1] J. A. Fox *et al.*, "Articular cartilage thickness of the humeral head: an anatomic study," *Orthopedics,* vol. 31, no. 3, p. 216, Mar 2008.

[2] C. Schleich, B. Bittersohl, G. Antoch, R. Krauspe, C. Zilkens, and J. Kircher, "Thickness Distribution of Glenohumeral Joint Cartilage: A Normal Value Study on Asymptomatic Volunteers Using 3-Tesla Magnetic Resonance Tomography," (in English), *Cartilage,* vol. 8, no. 2, pp. 105-111, Apr 2017.

[3] B. Couteau, P. Mansat, E. Estivalezes, R. Darmana, M. Mansat, and J. Egan, "Finite element analysis of the mechanical behavior of a scapula implanted with a glenoid prosthesis," *Clin Biomech (Bristol, Avon),* vol. 16, no. 7, pp. 566-75, Aug 2001.

[4] P. Clavert, M. Zerah, J. Krier, P. Mille, J. F. Kempf, and J. L. Kahn, "Finite element analysis of the strain distribution in the humeral head tubercles during abduction: comparison of young and osteoporotic bone," *Surg Radiol Anat,* vol. 28, no. 6, pp. 581-7, Dec 2006.

[5] P. Favre, P. Kloen, D. L. Helfet, and C. M. Werner, "Superior versus anteroinferior plating of the clavicle: a finite element study," *J Orthop Trauma,* vol. 25, no. 11, pp. 661-5, Nov 2011.

[6] M.-L. Yeh, D. Lintner, and Z.-P. Luo, "Stress distribution in the superior labrum during throwing motion," *The American journal of sports medicine,* vol. 33, no. 3, pp. 395-401, 2005.

[7] K. Sung-Woo, J. M. Cavanaugh, J. P. Leach, and S. W. Rouhana, "Mechanical properties of the shoulder ligaments under dynamic loading," *Stapp Car Crash Journal,* vol. 48, p. 125, 2004.

[8] C.-Y. Huang, A. Stankiewicz, G. A. Ateshian, and V. C. Mow, "Anisotropy, inhomogeneity, and tension–compression nonlinearity of human glenohumeral cartilage in finite deformation," *Journal of biomechanics,* vol. 38, no. 4, pp. 799-809, 2005.

[9] P. Buchler, N. A. Ramaniraka, L. R. Rakotomanana, J. P. Iannotti, and A. Farron, "A finite element model of the shoulder: application to the comparison of normal and osteoarthritic joints," (in English), *Clinical Biomechanics,* vol. 17, no. 9-10, pp. 630-639, Nov-Dec 2002.

[10] C. D. Smith, S. D. Masouros, A. M. Hill, A. L. Wallace, A. A. Amis, and A. M. J. Bull, "The Compressive Behavior of the Human Glenoid Labrum May Explain the Common Patterns of SLAP Lesions," (in English), *Arthroscopy-the Journal of Arthroscopic and Related Surgery,* vol. 25, no. 5, pp. 504-509, May 2009.

[11] G. Bergmann, F. Graichen, A. Bender, M. Kaab, A. Rohlmann, and P. Westerhoff, "In vivo glenohumeral contact forces--measurements in the first patient 7 months postoperatively," *J Biomech,* vol. 40, no. 10, pp. 2139-49, 2007.

[12] G. Bergmann *et al.*, "In vivo gleno-humeral joint loads during forward flexion and abduction," *J Biomech,* vol. 44, no. 8, pp. 1543-52, May 17 2011.

[13] F. C. Van der Helm, "Analysis of the kinematic and dynamic behavior of the shoulder mechanism," *Journal of biomechanics,* vol. 27, no. 5, pp. 527-550, 1994.

[14] A. Terrier, A. Vogel, M. Capezzali, and A. Farron, "An algorithm to allow humerus translation in the indeterminate problem of shoulder abduction," *Med Eng Phys,* vol. 30, no. 6, pp. 710-6, Jul 2008.

[15] P. Favre, M. Senteler, J. Hipp, S. Scherrer, C. Gerber, and J. G. Snedeker, "An integrated model of active glenohumeral stability," *J Biomech,* vol. 45, no. 13, pp. 2248-55, Aug 31 2012.

[16] L. Sins, P. Tétreault, N. Hagemeister, and N. Nuño, "Adaptation of the AnyBody™ musculoskeletal shoulder model to the nonconforming total shoulder arthroplasty context," *Journal of biomechanical engineering,* vol. 137, no. 10, 2015.

[17] M. X. Zheng, Z. H. Qian, Z. M. Zou, C. Peach, and L. Ren, "Subject-Specific Finite Element Modeling of the Human Shoulder Complex Part 2: Quantitative Evaluation of the Effect of Rotator Cuff Tear Propagation on Glenohumeral Joint Stability," (in English), *Ieee Access,* vol. 7, pp. 34068-34077, 2019.